## Sheet1
Vestibular equipment: Air Caloric Stimulator
Description
A device that delivers air to the ear canal at a controlled rate of flow and temperature to assess vestibular functioning of the patient’s balance system
Technical specifications
Temperature range: 12° to 50°C
Time to temperature < 60 seconds
Temperature readout: Digital
Temperature accuracy: ± 0.3°C
Time range: 1 - 99 seconds
Flow rate range: 4 - 10 LPM
Capability to control by computer: Standard
Input power: 120/240 VAC @ 50/60 Hz
5 A @ 120 VAC, 2.5 A @ 240 VAC (max.)
Safety features: Primary temperature read-out
Separate over-temperature cut-off
Stimulus timing activation: Switch on delivery head or footswitch
Size (W x D x H): 35 x 44 x 20 cm
Weight: 20kg
Specula: 8-62-44100 (pkg-25)
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list