## Sheet1
ITEM | X-RAY BUCKY UNIT WITH A FREE - STANDING TUBE STAND and VERTICAL BUCKY STAND.
1
1.1 | The said Unit must be fully automatic, digital, radiographic examination and evaluation workstation based on two detector technology for high image dynamics with excellent signal/noise ratio is required.
1.2 | The bid price quoted must include the supply, installation, demonstration and commissioning of the equipment as well as the building alterations that may be required for the successful installation of the unit/s.
1.3 | State the make and model of the unit offered.
1.4 | The bidders are requested to structure their quote so that it reflects the price of the unit separately from any building alterations that are required. The price schedule at the end of this specification must be completed.
1.5 | The unit must be suitable for general and trauma radiography.
TECHNICAL SPECIFICATION
2 | GENERATOR:
2.1 | The unit should be a high frequency type generator.
2.2 | The output must be 60kW. State maximum milliamps at different kilovolts
2.3 | The unit must incorporate an electronic timer with a minimum exposure time of 0,001 sec or less and electronic switching.
2.4 | The unit must include automatic overload protection.
2.5 | The unit must include automatic mains compensation.
2.6 | A milliamp meter or indicator must be provided
2.7 | The kV, mA and time must be independently selectable and digitally displayed.
2.8 | Bidders to state the range of the exposure timer. Acceptable range 0.001 sec – 6 sec (1 ms – 6sec).
2.9 | Bidders to state kV and mA ranges for radiography. The kV range to be at least 40 to 125 kV and mA range at least 10 to 600mA. KV selection must be in steps of 1 and must be a single control.
2.10. | The unit must preferably be programmable. Supply details of the programming facility and the number of programmes.
2.11 | The offered unit must include automatic exposure control.
3 | ROTATING ANODE X‑RAY TUBE
3.1 | The rotating anode tube must match the output of the generator.
3.2 | Bidders to supply the following details:
3.2.1 | The maximum tube kV must be at least 125 kV
3.2.2 | The maximum tube mA must be at least 600 mA
3.2.3 | The speed of anode rotation must be at least 8000rpm.
3.2.4 | Anode heat storage capacity must be at least 300kHU
3.2.5 | The tube housing heat storage capacity must be more than 400kHU
3.2.6 | The size of small focus must be 0,6 mm
3.2.7 | The size of large focus must be 1,2 mm
3.2.8 | Inherent filtration of at least 0.8mmAl
3.2.9 | Total filtration must exceed 2.5mmAl.
3.2.10 | Anode angle must be large enough to provide a uniform beam over a field of 43cm x 43cm at a FFD of 100cm.
3.3 | One pair of high tension cables of suitable length must be supplied. The cable length must not restrict free movement of the tube for all radiographic procedures.
3.4 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
4 | COLLIMATOR.
4.1 | The unit must have a multileaf collimator.
4.2 | Rotating flange between the collimator and tube must be supplied to provide diagonal collimation.
4.3 | One sinus cone must be supplied with the unit.
4.4 | A suitable handle bar with central release of all brakes as well as individual release must be provided.
4.6 | An electronic timer for the Halogen light source must be provided. The collimator to provide illumination in excess of 100 lux.
5 | TUBE STAND.
5.1 | A floor mounted tube stand must be supplied that will provide a focal film distance of at least 150 cms to the Bucky table.
The tube stand must not be integrated into the table.
The longitudinal movement of the tube stand must extend beyond the table at both ends.
5.2 | The unit must be fitted with electro‑magnetic brakes.
5.3 | Bidders to indicate the focal film distance range. Markings indicating the focal film distance on the tube stand must be provided.
5.4 | A safety device for possible failure of the counter‑balance system must be incorporated.
5.5 | The unit must be suitable to function with an x‑ray Bucky table, the vertical Bucky and provide a facility for horizontal beam radiography.
5.6 | Bidders to state:
i) Tube rotation around vertical axis
ii) Tube rotation around horizontal axis and
iii) Range of vertical movement.
iv) Range of longitudinal movement.
5.7 | Radiography of a patient who is in a seated position on the stretcher or table top must be possible.
5.8 | The unit must be suitable for radiography of patients on trolleys and therefore space must be available at both ends of the table – e.g. it must be possible to do a supine chest x-ray or a tibia and fibula x-ray on a hospital trolley. State the longitudinal movement of the tube column.
5.8 | Rotation of the tube to the lateral position for horizontal shoot through radiographs must be possible by means of a central electronic locking/unlocking device at the collimator. Units that have the locking/unlocking device on the tube stand close to the tube rail will not be accepted.
6 | BUCKY TABLE
6.1 | The Bucky table must have a flat, four way floating top.
6.2 | The table must be electronically height controlled. Foot switches must form part of the table.
6.3 | The bidder to state dimensions of table top offered (length and width).
6.4 | State the variable height of the table offered.
6.5 | State the lateral and longitudinal movement of the table top.
6.6 | The table offered must be able to carry patients with a weight of at least 220kg.
6.7 | State longitudinal movement of Bucky tray. Bidders must detail whether the distance relates to the outermost end of the Bucky tray or from centre point to centre point. A data sheet must be included to clarify this point.
6.8 | The table‑top movement must be controlled with electro‑magnetic brakes.
6.9 | The table‑top ‑ film distance must be less than 6 cms.
6.10. | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology.
6.11 | A compression band with fitting mechanism compatible to the top must be supplied.
6.12 | A full set of positioning sponges with washable water resistance covers must be supplied with the table.
7 | VERTICAL BUCKY
7.1 | State the make and model of the vertical Bucky offered.
7.2 | The Bucky assembly must be fitted with a oscillating secondary radiation grid of 10: 1, and be suitable for chest in Bucky radiography.
7.3 | Radiography of a patient who is in a seated position must be possible.
7.4 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology and must accommodate a chest radiography at 180cm.
7.5 | The Bucky must be mounted centrally on the column with the choice of loading on either left or right
7.6 | Vertical movement of the vertical Bucky must be via an electronic locking/unlocking switch. Analogue locking/unlocking levers or knobs will not be accepted
7.7 | An overhang frame for out of Bucky radiography must be included that is compatible with erect Bucky.
7.8 | An overhang arm support must be included that is compatible with erect Bucky.
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
8.5 | A radiation report must be automatically generated and transferable via DICOM.
8.6 | Should be FDA or CE approved product
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan